## BSC CERT DATABASE-2015
[7 empty] | Allowable Ranges |  | Intial measurements |  | Final Setpoints |  | HEPA Filter test
Facility name | County | Date of Servicing/Certification | Cabinet Type | Manufacturer | Model | Serial # | Intake | Downflow | Intake | Downflow | Intake | Downflow | Supply | Exhaust | Smoke Test | Ever serviced | Overall Outcome | Users trained | Date of next service | Overall Recommendation
Naivasha District Hospital | Nakuru | 23/6/15 | BSC class 11 Type A2 | Esco | AC2-4E1 | 2011-6345 | 84FPM-94FPM | 53-71FPM |  |  | 89 | 55 | OK | OK | Pass | no | Pass | 17 | 2016-06-01
Kiambu District Hospital | Kiambu | 24/6/15 | BSC class 11 Type A2 | Nuve | MN 090 | 2.0773 | 63FPM-95FPM | 73-87FPM |  |  | 79 | 75 | OK | OK | Pass | no | Pass | 13 | 2016-06-01
Gatundu District Hospital | Kiambu | 25/6/15 | BSC class 11 Type A2 | Labex | BSC 700IIA2 | TXA200111 | ≥98FPM | 53-71FPM |  |  | 101FPM | 64FPM | OK | OK | Pass | no | Pass | 7 | 2016-06-01
Gatundu District Hospital | Kiambu | 25/6/15 | Tb safety hood |  |  |  | ≥100FPM | [7 empty] | Pass | no | Pass | 7 | 2016-06-01
Muranga District Hospital | Muranga | 26/6/15 | BSC class 11 Type A2 | Labex | BSC-700IIA2 | TXA 200275 | ≥98FPM | 53-71FPM |  |  | 104FPM | 72FPM | OK | OK | Pass | yes | Pass | 12 | 2016-06-01
Muranga District Hospital | Muranga | 26/6/15 | Tb safety hood |  |  |  | ≥100FPM | [7 empty] | Pass | no | Pass | 12 | 2016-06-01
Nyeri Provincial General Hospital | Nyeri | 28/6/15 | BSC class 11 Type A2 | Thermo Scientific | 1376 Rel #1 | 110236-93 | 100-110 | 58-68 |  |  | 107FPM | 61FPM | OK | OK | Pass | yes | Pass | 6 | 2016-06-01
Nyeri Provincial General Hospital | Nyeri | 28/6/15 | BSC class 11 Type A2 | Esco | AC2-4E1 | 2009-42980 | 84FPM-94FPM | 53-71FPM |  |  | 91FPM | 57FPM | OK | OK | Pass | yes | Pass | 6 | 2016-06-01
Mt. Kenya Hospital -Nyeri | Nyeri | 29/6/15 | Tb safety hood |  |  |  | ≥100FPM | [7 empty] | Pass | no | Pass | 0 | 2016-06-01
Embu Provincial General Hospital | Embu | 30/6/15 | BSC class 11 Type A2 | Thermo Scientific | 1376 | 110236-87 | 100-110 | 58-68 |  | 56 | 102 | 62 | OK | OK | Pass | yes | Pass | 2 | 2016-06-01
Chuka District Hospital | Meru | 2015-01-07 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 222FPM |  |  |  | Pass | no | Pass | 14 | 2016-07-01
Miathene Sub County Hospital | Meru | 2015-02-07 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 113FPM |  |  |  | Pass | no | Pass | 3 | 2016-07-01
Mbeu Sub District Hospital | Meru | 2015-02-07 | Tb safety hood |  |  |  | ≥100FPM |  | 18FPM |  | 18FPM |  |  |  | Failed | no | Failed | 0 | ASAP
Nyambene Sub District Hospital | Meru | 2015-03-07 | Tb safety hood |  |  |  | ≥100FPM |  | 77FPM |  | 77FPM |  |  |  | Pass | no | Pass | 4 | 2016-07-01
Meru District Hospital | Meru | 2015-03-07 | Tb safety hood |  |  |  | ≥100FPM |  | 117FPM |  | 117FPM |  |  |  | Pass | no | Pass | 5 | 2016-07-01
Mwingi District Hospital | Kitui | 2015-06-07 | BSC class 11 Type A2 | Labex | BS8 | BSC 31A0381 | 84- 94FPM | 53- 71FPM |  |  | 85FPM | 56FPM | OK | OK | Pass | no | Pass | 6
Kitui District Hospital | Kitui | 2015-07-07 | BSC class 11 Type A2 | Labex | BS8 | BSC 31A0406 | ≥98FPM | 53- 71FPM |  |  | 106FPM | 58FPM | OK | OK | Pass | yes | Pass | 5 | 2016-07-01
Kitui KMTC | Kitui | 2015-08-06 | BSC class 11 Type A2 | Esco | AC2-3E1 | 2005-11732 | 84- 94FPM | 53- 71FPM |  |  | 89FPM | 66FPM | OK | OK | Pass | yes | Pass | 4 | 2016-07-01
Matuu District | Machakos | 2015-09-07 | BSC class 11 Type A2 | Labex | BS8 | BSC31A0410 | 84- 94FPM | 53- 71FPM |  |  | 105FPM | 58FPM | OK | OK | Pass | no | Pass | 3 | 2016-07-01
Kangundo District Hospital | Machakos | 2015-10-07 | BSC class 11 Type A2 | Labex | BS8 | BSC31A0401 | 84- 94FPM | 53- 71FPM |  |  | 84FPM | 53FPM | OK | OK | Pass | no | Pass | 2 | 2016-07-01
Machakos Level 5 Hospital | Machakos | 2015-11-07 | BSC class 11 Type A2 | Thermo Scientific | 1376 Rel # 1 | 110236-98 | 100-110 | 58-68 |  |  | 102FPM | 63FPM | OK | OK | Pass | yes | Pass | 11 | 2016-07-01
Mbooni District Hospital | Makueni | 13/7/2015 | BSC class 11 Type A2 | Labex | BS8 | BSC31A5011338 | 84- 94FPM | 53- 71FPM |  |  | 90FPM | 59FPM | OK | OK | Pass | no | Pass | 5 | 2016-07-01
Makueni County Hospital | Makueni | 14/7/15 | BSC class 11 Type A2 | Labex | BS8 | BSC 31A00004 | 84- 94FPM | 53- 71FPM |  |  | 97FPM | 56FPM | OK | OK | Pass | no | Pass | 6 | 2016-07-01
Kimbwezi District Hospital | Makueni | 15/7/15 | BSC class 11 Type A2 | Labex | BS8 | BSC 31A0660 | 84- 94FPM | 53- 71FPM |  |  | 109FPM | 58FPM | OK | OK | Pass | no | Pass | 3 | 2016-07-01
Makindu District Hospital | Makueni | 15/7/15 | BSC class 11 Type A2 | Labex | BS8 | BSC 31A0395 | 84- 94FPM | 53- 71FPM |  |  | 92FPM | 56FPM | OK | OK | Pass | no | Pass | 9 | 2016-07-01
Sultan Hamud District Hospital | Makueni | 16/7/15 | BSC class 11 Type A2 | Labex | BS8 | BSC 31A01342 | 84- 94FPM | 53- 71FPM |  |  | 107FPM | 59FPM | OK | OK | Pass | no | Pass | 3 | 2016-07-01
Aids Healthcare Foundation of Kenya (AHF-K) Clinic | Makueni | 17/7/15 | BSC class 11 Type A2 | Labex | BBS-70011 | TXA200292 | ≥98FPM | 53-71FPM |  |  | 94FPM | 72FPM | OK | OK | Pass | no | Pass | 1 | 2016-07-01
Isiolo District Hospital | Isiolo | 22/08/15 | Tb safety hood |  |  |  | 150-200FPM |  |  |  | 153FPM |  |  | OK | Pass | no | pass | 1 | 2016-08-01
Sotik Health Centre | Bomet | 27/08/15 | BSC class 11 Type A2 | Labex | BSC 700IIA2 | TXA 200114 | ≥98FPM | ≥59FPM |  |  | 100FPM | 70FPM | OK | OK | pass | no | pass | 4 | 2016-08-01
Ndanai S/District Hospital | Bomet | 26/08/15 | BSC class 11 Type A2 | Nuve | MN 090 | 02-0695 Class EN | 63-95FPM | 59-93FPM |  |  | 80FPM | 74FPM | OK | OK | pass | no | pass | 8 | 2016-08-01
Kapkatet District Hosptal | Kericho | 28/08/15 | BSC class 11 Type A2 | Labex | BSC 1100 IIA2-X | DXA201490 | ≥98FPM | ≥59FPM |  |  | 100FPM | 61FPM | OK | OK | pass | yes | pass | 8 | 2016-08-01
Kapkatet District Hosptal | Kericho | 28/08/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 123FPM |  |  |  | pass | yes | pass | 8 | 2016-08-01
Tenwek Mission Hospital | Bomet | 24/08/15 | BSC class 11 Type A2 | Esco | AC2-4E1 | 2011-56066 | 84-94FPM | 53-71FPM |  |  | 85FPM | 54FPM | OK | OK | pass | no | pass | 9 | 2016-08-01
Tenwek Mission Hospital | Bomet | 24/08/15 | BSC class 1 | LABCONCO | 3980321 | 050233189 B | ≥75FPM |  |  |  | 82FPM |  |  | OK | pass | no | pass | 9 | 2016-08-01
Longisa District Hospital | Bomet | 25/08/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 102FPM |  |  |  | pass | yes | pass | 7 | 2016-08-01
Miteitei District Hospital | Nandi | 30/08/15 | BSC class 11 Type A2 | Labex | BSC 70011 | TXA 200280 | 84-94FPM | 53-68FPM |  |  | 94FPM | 55FPM | OK | OK | pass | yes | pass | 2 | 2016-08-01
Roret Disctrict Hospital | Kericho | 29/08/15 | BSC class 11 Type A2 | Labex | BSC 70011 A2 | TXA200113 | 84-94FPM | 53-68FPM |  |  | 87FPM | 55FPM | OK | OK | pass | yes | pass | 2 | 2016-08-01
Roret Disctrict Hospital | Kericho | 29/08/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 100FPM |  |  |  | pass | no | pass | 2 | 2016-08-01
Nandi Hills District Hospital | Nandi | 31/08/15 | BSC class 11 Type A2 | Labex | 70011 | TXA 200286 | 84-94FPM | 53-71FPM |  |  | 90FPM | 67FPM | OK | OK | pass | yes | pass | 8 | 2016-08-01
Nandi Hills District Hospital | Nandi | 31/08/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 1110FPM |  |  |  | pass | no | pass | 8 | 2016-08-01
Kapsabet District Hospital | Nandi | 2015-01-09 | BSC class 11 Type A2 | Labex | BSC 7000011 | TXA200286 | ≥98FPM | 53-71FPM |  |  | 98FPM | 53FPM | OK | OK | pass | yes | Pass | 12 | 2016-09-01
St. Clares Kaplong Mission Hospital | Bomet | 2015-01-09 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 161FPM |  |  |  | pass | no | pass | 2 | 2016-09-01
Kabaranet Referral Hospital | Baringo | 2015-02-09 | BSC class 11 Type A2 | Baker | SG-303-INT | 113769 | 100-110FPM | 50-60FPM |  |  | 107FPM | 52FPM | OK | OK | pass | yes | pass | 6 | 2016-09-01
Kabaranet RefferalHospital | Baringo | 2015-02-09 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 106FPM |  |  |  | pass | yes | pass | 6 | 2016-09-01
Nakuru G.K Prisons | Nakuru | 2015-03-09 | BSC class 11 Type A2 | Labex | BS 8 | BSC 31A0400 | 84-94FPM | 53-71FPM |  |  | 92FPM | 53FPM | OK | OK | pass | yes | pass | 2 | 2016-09-01
Naivasha Maximum Prisons | Nakuru | 2015-03-09 | Tb safety hood |  |  |  | 100-110FPM |  |  |  | 204FPM |  |  |  | pass | no | pass | 2 | 2016-09-01
Mbagathi District Hospital | Nairobi | 18/03/15 | BSC class 11 Type A2 | Thermo Scientific | 1376 Rel # 1 | 110236-94 | 100-110FPM | 58-68FPM |  |  | 103FPM` | 60FPM | OK | OK | pass | yes | pass | 4 | 2016-03-01
Thika Level 5 Hospital | Kiambu | 28/05/2015 | BSC class 11 Type A2 | Thermo Scientific | 1378 Rel # 1 | 110236-92 | 100-110FPM | 58-68FPM |  |  | 102FPM | 59PFM | OK | OK | pass | yes | pass | 7 | 2016-05-01
Karatina District Hospital | Nyeri | 29/05/15 | BSC class 11 Type A2 | Labex | BSC 700IIA2-Z | TXA 200028 | ≥98FPM | ≥59FPM |  |  | 109FPM | 70FPM | OK | OK | pass | yes | pass | 2 | 2016-05-01
Gathigiriri Health Centre | Kirinyaga | 30/05/2015 | BSC class 11 Type A2 | Esco | AC2-4E8 | 2013-85709 | 84-94FPM | 53-71FPM |  |  | 91FPM | 68FPM | OK | OK | pass | no | pass | 1 | 2016-05-01
Kerugoya District Hospital | Kirinyaga | 31/05/15 | BSC class 11 Type A2 | Labex | BSC 700IIA2 | TXA 200020 | ≥98FPM | ≥59FPM |  |  | 114FPM | 78FPM | OK | OK | pass | yes | pass | 4 | 2016-05-01
Kimbimbi Sub District Hospital | Kirinyaga | 31/05/15 | BSC class 11 Type A2 | Nuve | MN 090 | 02-0605 | 63-94FPM | 47-71FPM |  |  | 86FPM | 68FPM | OK | OK | pass | no | pass | 3 | 2016-05-01
Malindi Sub County Hospital | Kilifi | 15/01/15 | BSC class 11 Type A2 | Thermo Scientific | 1378 Rel # 1 | 118491-788 | 100-110FPM | 58-68FPM |  |  | 107FPM | 61FPM | OK | OK | pass | yes | Pass | 8 | 2016-01-01
Huruma Lions Health Centre | Nairobi | Feb, 2015 | BSC class 11 Type A2 | Faster | EN 12469 | 149 | ≥98FPM | ≥59FPM |  |  | 18FPM | 10FPM |  |  | Failed | yes | failed | 0 | ASAP
Ngara Clinic | Nairobi | Feb, 2015 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 106FPM |  |  |  | pass | no | pass | 5 | 2016-02-01
Kibera D.O Clinic | Nairobi | Feb, 2015 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 19FPM |  |  |  | failed | no | failed | 0 | ASAP
Karen Health Centre | Nairobi | 2015-02-01 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 62FPM |  |  |  | Average | no | average | 3 | ASAP
Waithaka Health Centre | Nairobi | 2015-02-01 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 109FPM |  |  |  | pass | no | pass | 2 | 2016-02-01
Eastlegh Health Centre | Nairobi | Feb,2015 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 118FPM |  |  |  | pass | no | pass | 2 | 2016-02-01
Umoja Health Centre | Nairobi | 2015-02-01 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 200FPM |  |  |  | pass | no | pass | 2 | 2016-02-01
Chandaria Health Centre | Nairobi | 2015-02-01 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 91FPM |  |  |  | pass | no | pass | 2 | 2016-02-01
Embakasi Health Centre | Nairobi | 2015-02-01 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 230FPM |  |  |  | pass | no | pass | 1 | 2016-02-01
Dandora Health Centre | Nairobi | 2015-02-01 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 38FPM |  |  |  | failed | no | failed | 2 | ASAP
LOCO Health Centre | Nairobi | 2015-02-01 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 51FPM |  |  |  | failed | no | failed | 1 | ASAP
HCCK Health Centre | Nairobi | 2015-02-01 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 224FPM |  |  |  | pass | no | passed | 2 | 2016-02-01
Dagoreti Sub District Hospital | Nairobi | 2015-02-01 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 104FPM |  |  |  | pass | no | passed | 2 | 2016-02-01
Ruiru G.K Prisons | Kiambu | 2015-02-01 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 54FPM |  |  |  | failed | no | failed | 1 | ASAP
Kamiti Maximum Prisons | Kiambu | 2015-02-01 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 56FPM |  |  |  | failed | yes | failed | 2 | ASAP
Remand Prisons Health Centre | Nairobi | 2015-02-01 | BSC class 11 Type A2 | Labex | BS 8 | BSC 31 A0382 | 95-101FPM | 56-62FPM |  |  | 99FPM | 56FPM | OK | OK | pass | no | pass | 2 | 2016-02-01
Nairobi West Prisons Health Centre | Nairobi | 2015-02-01 | BSC class 11 Type A2 | Labex | BSC-700IIA2 | TXA 200108 | 95-101FPM | 56-62FPM |  |  | 97FPM | 59FPM | OK | OK | pass | no | pass | 1 | 2016-02-01
Langata Women Prisons Heath Centre | Nairobi | 2015-02-01 | BSC class 11 Type A2 | Esco | AC2-2E1 | 2010-46012 | 84-94FPM | 53-71FPM |  |  | 87FPM | 55FPM | OK | OK | pass | yes | pass | 2 | 2016-02-01
SOS Buruburu Clinic | Nairobi | 2015-02-01 | BSC class 11 Type A2 | Esco | AC2-4E1 | 2006-15386 | 84-94FPM | 53-71FPM |  |  | 89FPM | 66FPM | OK | OK | pass | no | pass | 1 | 2016-02-01
Mathare Hospital | Nairobi | 2015-02-01 | BSC class 11 Type A2 | MRC | BSC-1311A2-X | BSC-13A0013 | 88-108FPM | 54-64FPM |  |  | 94FPM | 60FPM | OK | OK | pass | yes | pass | 1 | 2016-02-01
Riruta Health Centr | Nairobi | 2015-02-01 | BSC class 11 Type A2 | MRC | BSC-13A0035 | BSC-1300IIA267 | ≥98FPM | ≥59FPM |  |  | 118FPM | 72FPM | OK | OK | pass | yes | pass | 1 | 2016-02-01
LVCT Health Centre | Nairobi | 2015-02-01 | BSC class 11 Type A2 | MRC | BSC-711A2 | TXA 200315 | 84-94FPM | 53-71FPM |  |  | 89FPM | 61FPM | OK | OK | pass | yes | pass | 2 | 2016-02-01
Dream Clinic | Nairobi | 2015-12-01 | BSC class 11 Type A2 | Esco | : AC2-4S1 | 2007-27822 | 84-94FPM | 53-71FPM |  |  | 91FPM | 60FPM | OK | OK | pass | no | pass | 6 | 2015-12-01
Dream Clinic | Nairobi | 2015-12-01 | BSC class 11 Type A2 | Thermo Scientific | MSC-Advantage | 40851530 | 100-110FPM | 58-68FPM |  |  | 106FPM | 63FPM | OK | OK | pass | no | pass | 6 | 2015-12-01
Mtongwe Dispensary | Mombasa | 14/09/15 | BSC class 11 Type A2 | Nuve | MN 090 | 20286 | ≥79FPM | 63-94FPM |  |  | 114FPM | 84FPM | OK | OK | pass | no | pass | 3 | 2016-09-01
Kenya Navy | Mombasa | 14/09/15 | BSC class 11 Type A2 | Esco | AC2-4E1 | 2007-20701 | 84-94FPM | 53-71FPM |  |  | 90FPM | 56FPM | OK | OK | pass | yes | pass | 2 | 2016-09-01
Bomu Clinic-Likoni | Mombasa | 15/09/15 | BSC class 11 Type A2 | Esco | AC2-4E1 | 2009-42985 | 84-94FPM | 53-71FPM |  |  | 89FPM | 58FPM | OK | OK | pass | no | pass | 2 | 2016-09-01
Mbuta Model Clinic | Mombasa | 15/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 143FPM |  |  |  | pass | no | pass | 2 | 2016-09-01
Shika Adabu Dispensary | Mombasa | 15/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 115FPM |  |  |  | pass | no | pass | 2 | 2016-09-01
King'orani Prisons Dispensary | Mombasa | 16/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 58FPM |  |  |  | failed | no | Failed | 2 | ASAP
Ganjoni Health Centre | Mombasa | 16/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 97fFPM |  |  |  | pass | no | pass | 3 | 2016-09-01
Kader Bhoy Dispensary | Mombasa | 16/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 108FPM |  |  |  | pass | no | pass | 4 | 2016-09-01
Shimo La Tewa Prisons | Mombasa | 17/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 139FPM |  |  |  | pass | no | pass | 3 | 2016-09-01
Utange Health Centre | Mombasa | 17/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | Nil |  |  |  | failed | no | failed(No electrical cables) |  | ASAP
Kongowea Health Centre | Mombasa | 17/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 126FPM |  |  |  | pass | no | pass | 2 | 2016-09-01
Likoni Sub County Hospital | Mombasa | 18/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 123FPM |  |  |  | pass | no | pass | 5 | 2016-09-01
Railway Clinic | Mombasa | 18/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 87FPM |  |  |  | pass | no | pass | 1 | 2016-09-01
Jocham Hospital | Mombasa | 18/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 82FPM |  |  |  | pass | no | pass | 5 | 2016-09-01
Aids Healthcare Foundation of Kenya (AHF-K) Clinic | Mombasa | 19/09/15 | BSC class 11 Type A2 | Labex | BBS-70011 | TXA200295 | 89-107FPM | 62-68FPM |  |  | 94FPM | 65FPM | OK | OK | pass | no | pass | 1 | 2016-09-01
Coast PGH | Mombasa | 14/09/15 | BSC class 11 Type A2 | Faster | BH-EN 2003 | BH-EN 200-S 341 | [9 empty] | yes |  | 4 | 2016-09-01
Bomu Hospital Research laboratory | Mombasa | 14/09/15 | BSC class 11 Type A2 | Esco | SC2-4A1 | 2008-34895 | [7 empty] | OK | pass | no | pass |  | 2016-09-01
Chaani Dispensary | Mombasa | 15/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 120FPM |  |  |  | pass | no | pass | 2 | 2016-09-01
Mikindani Catholic Dispensary | Mombasa | 15/09/15 | BSC class 11 Type A2 | Nuve | MN 090 | 02-123 | 84-94FPM | 53-71FPM |  |  | 92FPM | 59FPM | OK | OK | pass | yes | pass | 2 | 2016-09-01
Mikindani Municipal Diospensary | Mombasa | 15/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 49FPM |  |  |  | failed | no | failed | 2 | ASAP
Bomu Mariakani Hospital | Mombasa | 16/09/15 | BSC class 11 Type A2 | Esco | AC2-4E1 | 2011-56064 | 84-94FPM | 53-71FPM |  |  |  |  | OK | OK | pass | no | pass | 1 | 2016-09-01
Coast PGH | Mombasa | 14/09/15 | PCR | FISHER SCIENTIFIC |  | 100222018 | [6 empty] | OK | OK | pass | no | pass | 4 | 2016-09-01
Coast PGH | Mombasa | 14/09/15 | BSC class 11 Type A2 |  |  | W001896 | 89-99FPM |  |  |  | 79FPM |  | NOT OK |  | failed | yes | failed | 4 | ASAP
Coast PGH | Mombasa | 15/09/15 | Laminah Flow |  | SLEE | [7 empty] | OK | OK | pass | yes | pass | 4 | 2016-09-01
Tudor Sub District Hospital | Mombasa | 15/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | Nil |  |  |  | Failed | no | fail (Stuck motor) | 0 | ASAP
KPA Bandari Dispensary | Mombasa | 16/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 80FPM |  |  |  | pass | no | average | 3 | 2016-09-01
Mvita Sub County Hospital | Mombasa | 16/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 45FPM |  |  |  | failed | no | failed | 2 | ASAP
Bokole Health Centre | Mombasa | 16/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 137FPM |  |  |  | pass | no | pass | 3 | 2016-09-01
Jomvu Mikindani | Mombasa | 17/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 37FPM |  |  |  | failed | no | failed |  | ASAP
Miritini CDF Dispensary | Mombasa | 17/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 123FPM |  |  |  | pass | no | pass | 3 | 2016-09-01
Magongo Dispensary | Mombasa | 17/09/15 | Tb safety hood |  |  |  | ≥100FPM |  |  |  | 189FPM |  |  |  | pass | no | pass | 2 | 2016-09-01
Port Reirz | Mombasa | 18/09/15 | BSC class 11 Type A2 | BECTON DICKINSON | 60631 | 84020799 | 79-89FPM |  |  |  | 80FPM |  | OK | OK | pass | no | pass | 3 | 2016-09-01
Bomu Hospital Magongo | Mombasa | 18/09/15 | BSC class 11 Type A2 | Esco | AC2-4E1 | 2009-42979 | 84-94FPM | 53-71FPM |  |  | 86FPM | 59FPM | OK | OK | pass | no | pass | 7 | 2016-09-01
Mewa Hospital | Mombasa | 19/09/15 | BSC class 11 Type A2 | Esco | AC2-4E1 | 2009-38179 | 84-94FPM | 53-71FPM |  |  | 91FPM | 61FPM | OK | OK | pass | no | pass | 3 | 2016-09-01
AIC Liten Mission Hospital | Kericho | 2015-03-11 | BSC class 11 Type A2 | Esco | AC2-4E1 | 2007-23752 | 84-94FPM | 53-71FPM |  |  | 86FPM | 56FPM | OK | OK | pass | no | pass | 11 | 2016-11-01
Uniliver Central Hospital | Kericho | 2015-05-11 | BSC class 1 | LABCONCO | 3980321 | 050233192 B | ≥75FPM |  |  |  | 76FPM |  |  | OK | pass | no | pass | 3 | 2016-11-01
AIC Kijabe Naivasha Medical Clinic | Nakuru | 2015-07-11 | BSC class 1 | Karstylan Mettallyor | ANTIB 550 | 3020/06 | ≥75FPM |  |  |  | 48FPM |  |  | OK | pass | no | average | 2 | ASAP
Beacon of Hope Health Centre | Kajiado | 2015-09-11 | BSC class 11 Type A2 | Labex | BSC 700II | TXA 200017 | 84-94FPM | ≥59FPM |  |  | 87FPM | 74FPM | OK | OK | pass | no | pass | 1 | 2016-11-01
Nakuru county referral hospital | Nakuru | 2015-05-01 | BSC class 11 Type A2 | Thermo Scientific | 1376 Rel#1 | 110236 - 87 | 100-110FPM | 58-68FPM | 107FPM | 55FPM | 107FPM | 62FPM | OK | OK | pass | yes | pass | 0 | 2016-05-01
Olkalau Hospital | Nyandarua | 2015-05-01 | BSC class 11 Type A2 | Labex | 70011 | 1 A200299 | 89-107FPM | 62-68FPM | 99FPM | 68FPM | 99FPM | 68FPM | OK | OK | pass | yes | pass | 4 | 2016-05-01
Nyahururu County referral hospital | Nyandarua | 2015-05-01 | BSC class 11 Type A2 | Labex | BSC- 7001 | 1XA200279 | 89-107FPM | 62-68FPM | 101fpm | 62fpm | 101fpm | 62fpm | OK | OK | pass | yes | pass | 5 | 2016-05-01
Engineer County Hospital | Nyandarua | 2015-05-01 | BSC class 11 Type A2 | Labex | BSC - 700HA2 | TX A20015 | 89-107FPM | 62-68FPM | 100fpm. | 67fpm | 100fpm. | 67fpm | OK | OK | pass | yes | pass | 1 | 2016-05-01
Kitale County Hospital | Tran-nzoia | 2015-05-01 | BSC class 11 Type A2 | Thermo Scientific | 1376 REL 1 | 110236-84 | 100-110FPM | 58-68FPM | 107fpm. | 61fpm. | 107fpm. | 61fpm. | OK | OK | pass | yes | pass | 2 | 2016-05-01
Lodwar Countr Hospital | Turkana | 2014-10-01 | BSC class 11 Type A2 | Esco | AC2-4E1 |  | 84-94FPM | 53-71FPM |  |  |  |  | OK | OK | pass | no | pass | 0 | 2015-10-01 | additional AC unit is required
Kakuma refugee camp | Turkana | 2014-10-01 | BSC class 11 Type A2 | Esco | AC2-4E1 |  | 84-94FPM | 53-71FPM |  |  |  |  | OK | OK | pass | no | pass | 0 | 2015-10-01
Kakuma Mission Hospital | Turkana | 2014-10-01 | BSC class 11 Type A2 | Esco | AC2-4E1 |  | 84-94FPM | 53-71FPM |  |  |  |  | OK | OK | pass | no | pass | 2 | 2015-10-01
AIC Lokichogio | Turkana | 2014-10-01 | BSC class 1 | Chemfree Faster | 60-S | 10098 | [6 empty] | no | no | Average | yes | average |  | 2015-10-01 | Urgent replacement of HEPA filter is required
Lopiding Subcounty Hospital | Turkana | 2014-10-01 | BSC class 1 | ATIB | 1445/02 | [7 empty] | OK | OK | pass | no | pass |  | 2015-10-01
Lokitaung Subcounty Hospital | Turkana | 2014-10-01 | BSC class 1 | Chemfree Faster | 60-S | 201 | [6 empty] | OK | OK | pass | no | pass |  | 2015-10-01
Lokichar Subcounty Hospital | Turkana | 2014-10-01 | BSC class 1 | Chemfree Faster | 60-S | [7 empty] | OK | OK | pass | no | pass |  | 2015-10-01
TB Lab-NPHLS | Nairobi | 27/03/2015 | BSC class 11 Type A2 | Scanlaf | Mars Ptoros | L.11120934 | 100-110FMP | 65-75FPM |  |  | 105FPM | 71FPM | OK | OK | pass | yes | pass | 8 | 2016-03-01
TB Lab-NPHLS | Nairobi | 2015-09-11 | BSC class 11 Type A2 | Thermo Scientific | 1285 Rel#6 | 104830-1089 | 105-115FPM | 65-75FPM |  |  | 111FPM | 69FPM | OK | OK | pass | yes | pass | 8 | 2016-11-01
TB Lab-NPHLS | Nairobi | 2015-09-11 | BSC class 11 Type A2 | Thermo Scientific | 1376 Rel#1 | 110445-101 | 100-110FPM | 58-68FPM |  |  | 104FPM | 64FPM | OK | OK | pass | yes | pass | 8 | 2016-11-01
TB Lab-NPHLS | Nairobi | 2015-10-11 | BSC class 11 Type A2 | Thermo Scientific | 1376 Rel#1 | 110236-91 | 100-110FPM | 58-68FPM |  |  | 108FPM | 63FPM | OK | OK | pass | yes | pass | 8 | 2016-11-01
TB Lab-NPHLS | Nairobi | 2015-10-05 | BSC class 11 Type A2 | Baker | SG404-INT | 113160 | 100-110FPM | 58-68FPM |  |  | 106FPM | 60FPM | OK | OK | pass | new | pass | 8 | 2016-05-01
TB Lab-NPHLS | Nairobi | 2015-10-05 | BSC class 11 Type A2 | Baker | SG404-INT | 113145 | 100-110FPM | 58-68FPM |  |  | 103FPM` | 58FPM | OK | OK | pass | new | pass | 8 | 2016-05-01
TB Lab-NPHLS | Nairobi | 2015-10-05 | BSC class 11 Type A2 | Baker | SG404-INT | 113168 | 100-110FPM | 58-68FPM |  |  | 109FPM | 63FPM | OK | OK | pass | new | pass | 8 | 2016-05-01
TB Lab-NPHLS | Nairobi | 2015-10-05 | BSC class 11 Type A2 | Baker | SG404-INT | 113157 | 100-110FPM | 58-68FPM |  |  | 103FPM` | 61FPM | OK | OK | pass | new | pass | 8 | 2016-05-01
TB Lab-NPHLS | Nairobi | 2015-09-11 | BSC class 11 Type A2 | Lab Caire | SCI2R | SCR-197 | 100-110FPM | 58-68FPM |  |  | 106FPM | 63FPM | OK | OK | pass | yes | pass | 8 | 2016-11-01
NHRL Lab-NPHLS | Nairobi | 2015-09-05 | BSC class 11 Type A2 | Baker | SG404-INT | 113154 | 100-110FPM | 58-68FPM |  |  | 106FPM | 59FPM | OK | OK | pass | new | pass | 11 | 2016-05-01
NHRL Lab-NPHLS | Nairobi | 2015-09-05 | BSC class 11 Type A2 | Baker | SG404-INT | 113159 | 100-110FPM | 58-68FPM |  |  | 102FPM | 61FPM | OK | OK | pass | new | pass | 11 | 2016-05-01
NHRL Lab-NPHLS | Nairobi | 2015-08-05 | BSC class 11 Type A2 | Baker | SG404-INT | 113158 | 100-110FPM | 58-68FPM |  |  | 106FPM | 66FPM | OK | OK | pass | new | pass | 11 | 2016-05-01
NHRL Lab-NPHLS | Nairobi | 2015-08-05 | BSC class 11 Type A2 | Baker | SG404-INT | 113153 | 100-110FPM | 58-68FPM |  |  | 108FPM | 65FPM | OK | OK | pass | new | pass | 11 | 2016-05-01
NHRL Lab-NPHLS | Nairobi | 27/2/2015 | BSC class 11 Type A2 | Thermo Scientific | 1285 Rel#6 | 101651-314 | 105-115FPM | 65-75FPM |  |  | 108FPM` | 72FPM | OK | OK | pass | yes | pass | 11 | 2016-02-01
NHRL Lab-NPHLS | Nairobi | 20/1/2015 | BSC class 11 Type A2 | Thermo Scientific | 1285 Rel#6 | 101681-316 | 105-115FPM | 65-75FPM |  |  | 109FPM | 69FPM | OK | OK | pass | yes | pass | 11 | 2016-01-01
NHRL Lab-NPHLS | Nairobi | 20/1/2015 | BSC class 11 Type A2 | Thermo Scientific | 1285 Rel#6 | 101681-313 | 105-115FPM | 65-75FPM |  |  | 107FPM | 70FPM | OK | OK | pass | yes | pass | 11 | 2016-01-01
NHRL Lab-NPHLS | Nairobi | 21/1/2015 | BSC class 11 Type A2 | Thermo Scientific | 1285 Rel#6 | 101524-286 | 105-115FPM | 65-75FPM |  |  | 111FPM | 72FPM | OK | OK | pass | yes | pass | 11 | 2016-01-01
NHRL Lab-NPHLS | Nairobi | 21/1/2015 | BSC class 1 | Aura | Aura Mini | K03PO1N3718 | ≥75FPM |  |  |  | 81FPM |  |  | OK | pass | yes | pass | 11 | 2016-01-01
Influeza Lab-NPHLS | Nairobi | 23/7/2015 | BSC class 11 Type A2 | Thermo Scientific | 1388 Rel#4 | 117630-195 | 100-110FPM | 58-68FPM |  |  | 106FPM | 62FPM | OK | OK | pass | new | pass | 3 | 2016-07-01
Influeza Lab-NPHLS | Nairobi | 23/7/2015 | BSC class 11 Type A2 | Thermo Scientific | 1388 Rel#4 | 117630-198 | 100-110FPM | 58-68FPM |  |  | 103FPM` | 66FPM | OK | OK | pass | new | pass | 3 | 2016-07-01
Influeza Lab-NPHLS | Nairobi | 2015-09-09 | BSC class 11 Type A2 | Thermo Scientific | MSC-Advantage | 020041-01112 | 100-110FPM | 58-68FPM |  |  | 106FPM | 61FPM | OK | OK | pass | new | pass | 3 | 2016-09-01
Molecular Lab-NPHLS | Nairobi | 22/7/2015 | BSC class 11 Type A2 | Thermo Scientific | MSC-Advantage | 41198742 | 100-110FPM | 58-68FPM |  |  | 106FPM | 59FPM | OK | OK | pass | yes | pass | 5 | 2016-07-01
Microbiology Lab-NPHLS | Nairobi | 2015-12-02 | BSC class 11 Type A2 | Thermo Scientific | 1376 Rel#1 | 110236-85 | 100-110FPM | 58-68FPM |  |  | 107FPM | 65FPM | OK | OK | pass | yes | pass | 9 | 2016-02-01
Microbiology Lab-NPHLS | Nairobi | 24/2/2015 | BSC class 11 Type A2 | Thermo Scientific | 1386 Rel#4 | 117629-774 | 100-110FPM | 58-68FPM |  |  | 102FPM | 59FPM | OK | OK | pass | yes | pass | 9 | 2016-02-01
Microbiology Lab-NPHLS | Nairobi | 18/02/2015 | BSC class 11 Type A2 | Thermo Scientific | 1386 Rel#4 | 117694-784 | 100-110FPM | 58-68FPM |  |  | 107FPM | 64FPM | OK | OK | pass | yes | pass | 9 | 2016-02-01
Microbiology Lab-NPHLS | Nairobi | 24/2/2015 | BSC class 11 Type A2 | Baker | SG403AINT | 91590 | 100-110FPM | 58-68FPM |  |  | 106FPM | 62FPM | OK | OK | pass | new | pass | 9 | 2016-02-01
Spinal Injury Hospital | Nairobi | May, 2015 | BSC class 11 Type A2 | Esco | AC2-4E1 | 2011-65057 | 84-94FPM | 53-71FPM |  |  | 0FPM | 0FPM | NOT OK | NOT OK | Failed | yes | failed(blown out power supply) |  | ASAP

## BSC NOS
BSC previously certified limited to= 45
|  | BSC | TB Hoods | Totals
Current numbers of BSCs and TB hoods | 2014- 2015 | 237 | 86 | 323
Current numbers certified/serviced | 2015-10-01 | 121 | 57 | 178
|  |  | To be accounted for=NPHLS and Nyanza region | 51 BSCs and TB hoods
Previously the team used to certify only 45 BSCs

## Data summary
Cabinet type | No
BSC class I | 8
BSC class II | 94
TB Safety hoods | 50
PCR-Laminar flow | 2
Total | 154
ever serviced
yes | 54
no | 100
Total | 154
Smoke test performance
Pass | 137
average | 2
fail | 15
Total | 154
Status after calibration
Pass | 135
average | 4
fail | 15
Total | 154
missing data from the kemri engineer who worked on the western and nyanza region BSCs=33